# 1. Introduction

During their job firefighters are exposed to a high thermal load due to heavy physical activity, external heat exposure from fires and the wear of highly insulated protective clothing. High thermal load can cause heat stress     , resulting in heat exhaustion, dehydration, mental confusion, physical fatigue and loss of consciousness which affects productivity and risk perception. To monitor and prevent heat stress among firefighters, a reliable and continuous thermometer which is able to measure the real-time core temperature of firefighters is desirable.   .

Invasive core temperature (Tc) measurements may not be practical in a working situation The invasive temperature sensor pill is minimally invasive, but at the moment it is only available for remote Tc monitoring for a specific period of time and it is difficult to standardize the location of the sensor in the gastrointestinal tract  . Additionally, the pill must be swallowed at least 4–6 h prior to the measurement which is difficult in occupations such as firefighting because it is unknown when duty will call. Moreover, temperature sensor pills are impractical due to the high cost and inability for them to be reused  . In addition, food and liquid intake can influence the accuracy of the temperature sensor pill and higher body weights and/or abdominal proportions obstruct reading of the sensor. Other methods and research concerning non-obstructive measurement or prediction of Tc, e.g., via skin temperature or multiple parameters are not yet reliable or available for workers  .

The Cosinussᵒ C-med (Cosinussᵒ GmbH, München, Germany) is a new wearable, non-obstructive and commercially available inner-ear thermometer that could be useful to monitor Tc continuously and in a non-invasive manner. This sensor system could provide more detailed and long-term insight in the change in Tc during firefighting activities, as well as the role of different types of protective clothing, either as an individual measuring system or in combination with multiple variables In research of Chaglla  the Cosinuss One demonstrated a deviation of 1.5 C in comparison to inner-ear infrared (IR) thermometry. The correlation of the Cosinussᵒ C-med and compared to the research standard of gastrointestinal temperature is unknown     . The objective of this study is to investigate the validity of the Cosinuss to monitor changes in core temperature during realistic physical active firefighting simulation tasks instead of standard lab controlled treadmill protocols to mimic the real-life situation as well as possible. The aim of this study was (1) to test the validity and reliability of a wearable non-invasive Tc sensor, Cosinuss, in rest in comparison to an invasive temperature sensor pill and standard inner-ear IR thermometer and (2) during realistic firefighting simulation tasks in comparison to an invasive temperature sensor pill, and (3) to compare the change in Tc recorded with the Cosinuss and an invasive temperature sensor pill of firefighters during realistic firefighting simulation tasks in two types (traditional turnout gear versus a new concept) of protective clothing.

# 2. Materials and methods

1. معرفی

آتش نشانان در حین کار خود به دلیل فعالیت بدنی سنگین ، قرار گرفتن در معرض گرمای خارجی در برابر آتش سوزی و همچنین پوشیدن لباس محافظ بسیار عایق ، در معرض فشار زیاد حرارتی قرار دارند. بار حرارتی بالا می تواند باعث استرس گرما شود و در نتیجه فرسودگی گرما ، کمبود آب بدن ، سردرگمی ذهنی ، خستگی جسمی و از دست دادن هوشیاری که بر بهره وری و درک خطر تأثیر می گذارد ، باشد. برای نظارت و جلوگیری از استرس گرما در میان آتش نشانان ، یک دماسنج معتبر و مداوم که قادر به اندازه گیری دمای اصلی زمان واقعی آتش نشانان است ، مطلوب است. .

اندازه گیری دمای هسته تهاجمی (TC) ممکن است در شرایط کاری عملی نباشد قرص سنسور تهاجمی دما حداقل تهاجمی است ، اما در حال حاضر فقط برای یک کنترل از راه دور Tc برای یک دوره خاص در دسترس است و استاندارد سازی محل دشوار است سنسور در دستگاه گوارش. علاوه بر این ، قرص باید حداقل 4-6 ساعت قبل از اندازه گیری بلعیده شود که در مشاغل مانند آتش نشانی دشوار است زیرا نامشخص است هنگام انجام وظیفه چگونه است. علاوه بر این ، قرص های حسگر دما به دلیل هزینه بالا و عدم توانایی استفاده مجدد از آنها غیر عملی هستند. علاوه بر این ، مصرف مواد غذایی و مایع می تواند بر دقت قرص سنسور دما و وزن بالاتر بدن و / یا نسبت های شکمی مانع از خواندن سنسور شود. روش ها و تحقیقات دیگر در مورد اندازه گیری یا پیش بینی غیر انسدادی Tc ، به عنوان مثال ، از طریق درجه حرارت پوست یا پارامترهای متعدد هنوز قابل اعتماد یا در دسترس کارگران نیست.

Cosinussᵒ C-med (Cosinussᵒ GmbH ، مونشن ، آلمان) یک دماسنج جدید داخلی پوشیدنی ، بدون انسداد و تجاری در دسترس گوش داخلی است که می تواند برای نظارت بر Tc به طور مداوم و به روش غیر تهاجمی مفید باشد. این سیستم سنسور می تواند بینشی دقیق تر و طولانی مدت در مورد تغییر Tc در حین انجام فعالیت های آتش نشانی و همچنین نقش انواع مختلف لباس های محافظ ، چه به عنوان یک سیستم اندازه گیری فردی و چه در ترکیب با متغیرهای مختلف در تحقیقات Chaglla ارائه دهد. Cosinuss One در مقایسه با دماسنج مادون قرمز مادون قرمز گوش (IR) انحراف 1.5 درجه سانتیگراد را نشان داد. همبستگی Cosinussᵒ C-med و مقایسه آن با استاندارد تحقیقاتی دمای دستگاه گوارش ناشناخته است. هدف از این مطالعه ، بررسی اعتبار Cosinuss برای نظارت بر تغییرات در دمای هسته در طول کارهای واقعی شبیه سازی آتش نشانی فعال فیزیکی به جای پروتکل های تردمیل کنترل شده آزمایشگاهی استاندارد برای تقلید از وضعیت زندگی واقعی و همچنین ممکن است. هدف از این مطالعه (1) بررسی روایی و اعتبار یک سنسور غیر قابل حمل Tc پوشیدنی ، Cosinuss در حالت استراحت در مقایسه با یک قرص سنسور تهاجمی دما و دماسنج استاندارد IR داخلی گوش و (2) در حین آتش نشانی واقعی بود. وظایف شبیه سازی در مقایسه با یک قرص سنسور تهاجمی دما و (3) برای مقایسه تغییر در Tc ثبت شده با Cosinuss و یک قرص سنسور دمای تهاجمی آتش نشانان در حین انجام وظایف شبیه سازی آتش نشانی واقع گرایانه در دو نوع (دنده های مشارکت سنتی در مقابل یک مفهوم جدید) لباس محافظ

2. مواد و روش ها

## 2.1. Subjects

The subjects participated voluntarily and were recruited by distributing flyers and during an information meeting organized by the fire department via the local safety region. Inclusion criteria were firefighters with an age between 18 and 67 years who passed the Periodic Preventive Medical Examination (PPMO). Exclusion criteria were body weights lower than 40 kg, problems or complaints with the gastrointestinal tract and/or infestation of propreflex, and needing to undergo Nuclear Magnetic Resonance Imaging or Magnetic Resonance Imaging in the next 24 h after swallowing the thermometer pill (HQInc., 2018a, b). Subject information was protected by double-blinding the data; the fire department gave every subject a letter and the researchers coupled a number to this letter using a random number generator.

The Medical Ethics Committee of the University Medical Center Groningen, the Netherlands, issued a waiver for this study, stating that it does not involve medical research under Dutch law and approved the study (M17.209969). 2.2. Materials

### 2.2.1. Cosinussᵒ

The Cosinussᵒ type C-med (Cosinussᵒ GmbH, München, Germany) is a wearable core thermometer which measures the temperature in the inner-ear. This hearing-aid shaped thermometer can be used in working conditions of 15 to 55 C. According to Cosinuss (Cosinuss, 2016) the sample frequency is 100 Hz and the accuracy is 0.1 C.

### 2.2.2. CorTemp®

The CorTemp® HT150002 is an ingestible core body temperature sensor (dimensions: 2.4 10.7 mm) that has been approved by the US Food and Drug Administration (no. K880639)). This thermometer sensor pill has a temperature range 30 to 40ᵒC and an accuracy of 0.1ᵒC (). The pill data was continuously collected using the CorTemp® Data Recorder (dimensions

120x60 25mm, 193 g) with a sampling rate of 10 s The data recorder used CorTrack™ II Software version 2.7 () and needed to be worn around the hips.

### 2.2.3. Inner-ear infrared thermometer

The Braun ThermoScan® 7 type IRT 6520 (Braun GmbH, Kornberg, Germany) is an inner-ear IR thermometer. Due to its fast, easy to use and non-invasive nature, inner-ear IR thermometry is being used as a clinical standard. The Braun ThermoScan® has an measurement range temperature of 35 to 42 C with an accuracy of 0.2 C compared to rectal temperature measurements in an operating Ta of 10 to 40 C (Braun GmbH, sd)

### 2.2.4. Ambient conditions box

To measure the ambient temperature (Ta) and relative humidity (RH) (SHT15 Breakout, Sensirion, Staefa ZH, Switzerland) inside the protective clothing, an ambient conditions box was worn. The box was positioned on the chest using elastic belts and was worn over the first layer of clothing (a cotton t-shirt) and below the fire suit measuring the micro- climate inside the personal protective clothing. The temperature inside the clothing was described as Tcli (Lotens, 1993). The Tcli sensor has an accuracy of 0.3 C and the RH sensor has an accuracy of 2.0% at a range of 10–90% RH (Sensirion, 2010). The response time was 5–20 s and the operating temperature was 40 to 120 C (Sensirion, 2010).

2.1. فاعل، موضوع

افراد داوطلبانه شرکت کردند و با توزیع آگهی ها و طی یک جلسه اطلاعاتی که توسط سازمان آتش نشانی از طریق منطقه ایمنی محلی سازماندهی شده بود ، استخدام شدند. معیارهای ورود به مطالعه آتش نشانان با سن 18 تا 67 سال بودند که از معاینات پزشکی پیشگیری دوره ای (PPMO) گذشتند. معیارهای خروج از وزن ، وزن کمتر از 40 کیلوگرم ، مشکلات یا شکایات دستگاه گوارش و یا آلوده شدن پروپلکس و نیاز به تصویربرداری با رزونانس مغناطیسی هسته ای یا تصویربرداری با رزونانس مغناطیسی در 24 ساعت بعد از بلعیدن قرص دماسنج (HQInc.) 2018a ، ب). اطلاعات موضوع با دوبار کور کردن داده ها محافظت می شد. سازمان آتش نشانی نامه ای را به هر موضوعی ارائه می دهد و محققان با استفاده از تولید کننده عدد تصادفی عدد را با این نامه همراه می کنند.

کمیته اخلاق پزشکی مرکز پزشکی دانشگاه گرونینگن ، هلند ، با اعلام اینکه این موضوع طبق قانون هلند شامل تحقیقات پزشکی نمی شود ، از این تحقیق چشم پوشی کرد و این مطالعه را تأیید کرد (M17.209969). 2.2. مواد

2.2.1. Cosinussᵒ

Cosinussᵒ نوع C-med (Cosinussᵒ GmbH ، مونشن ، آلمان) یک دماسنج هسته پوشیدنی است که دما را در گوش داخلی اندازه گیری می کند. این دماسنج به شکل سمعک می تواند در شرایط کار 15 تا 55 درجه سانتیگراد استفاده شود. مطابق گفته Cosinuss (Cosinuss ، 2016) فرکانس نمونه 100 هرتز و دقت آن 0/0 درجه سانتیگراد است.

2.2.2. CorTemp®

CorTemp® HT150002 یک سنسور دمای بدن قابل هضم (ابعاد: 2.4 10.7 میلی متر) است که توسط سازمان غذا و داروی ایالات متحده تأیید شده است (شماره K880639). این قرص سنسور دماسنج دارای محدوده دما 30 تا 40 درجه سانتیگراد و دقت 0.1 درجه سانتیگراد () است. داده های قرص به طور مداوم با استفاده از ضبط کننده داده های CorTemp® (ابعاد) جمع آوری می شد

120x60 25 میلی متر ، 193 گرم) با نرخ نمونه گیری 10 ثانیه ضبط کننده اطلاعات از نرم افزار CorTrack ™ II نسخه 2.7 () استفاده کرده و در اطراف باسن پوشیده شده است.

2.2.3 دماسنج مادون قرمز گوش داخلی

Braun ThermoScan® 7 نوع IRT 6520 (Braun GmbH ، Kornberg ، آلمان) یک دماسنج IR داخلی گوش است. با توجه به ماهیت سریع ، آسان و غیر تهاجمی ، دماسنج IR داخلی گوش به عنوان یک استاندارد بالینی مورد استفاده قرار می گیرد. Braun ThermoScan® دارای دمای محدوده اندازه گیری از 35 تا 42 درجه سانتیگراد با دقت 0.2 درجه سانتیگراد در مقایسه با اندازه گیری دمای رکتوم در یک تائید عملیاتی از 10 تا 40 درجه سانتیگراد است (Braun GmbH، sd)

2.2.4 جعبه شرایط محیط

برای اندازه گیری دمای محیط (TA) و رطوبت نسبی (RH) (SHT15 Breakout ، Sensirion ، Staefa ZH ، سوئیس) در داخل لباس محافظ ، یک جعبه شرایط محیطی پوشیده شد. جعبه با استفاده از کمربندهای الاستیک روی قفسه سینه قرار گرفت و بیش از اولین لایه لباس (تی شرت پنبه ای) و در زیر لباس آتش اندازه گیری میزان آب و هوای داخل لباس محافظ شخصی پوشیده شد. درجه حرارت داخل لباس به عنوان Tcli توصیف شده است (Lotens، 1993). سنسور Tcli از دقت 0.3 درجه سانتیگراد و سنسور RH دارای دقت 2.0٪ در دامنه 10-90٪ RH است (Sensirion، 2010). زمان پاسخ 5-20 ثانیه و دمای کار 40 تا 120 درجه سانتیگراد بود (Sensirion، 2010).

### 2.2.5. Personal protective clothing

Two types of personal protective clothing were used; suit A and B. Suit A is traditional turnout gear composed of a trouser with jacket (see Fig. 1 left). Both trouser and jacket contained three layers; (1) outer fabric made of XT5 Nomex® Delta T; (2) moisture barrier made of MO3 Gore-tex® fireblocker; and (3) thermal barrier and inner layer made of Q01 thermal felt quilted to Nomex® viscose (standard EN469:2005) (Bristol Uniforms, 2007). This suit has a water vapor resistance of 28m2Pa/W, a thermal heat insulation (HTI) resistance HTI 24 of 21.3s and HTI 24-12 of 6.3s, and a radiation thermal heat insulation (RHTI) RHTI 24 of 26.2 s and RHTI 24-12 of 7.1s. Suit B is a new protective clothing concept composed of a coverall with jacket (see Fig. 1 right). The jacket contained two layers; (1) an outer fabric made of TenCate

Millenia™ Mi 9200 and (2) an inner layer made of TenCate Defender™ CZ 760. The coverall contains three layers; (1) an outer fabric made of TenCate Mllenia™ MI 9200 and Tecasafe® Plus XL 9700; (2) an inner layer of TenCate Defender™ CZ760 (upper part) and DM9180 (lower part); and (3) a thermal barrier of TenCate thermical membrane CX 140 (Safety Masters, 2017). This suit has a water vapor resistance of 12.6m2Pa/W, a thermal heat insulation (HTI) resistance HTI 24 of 17.4s and HTI 24-12 of 5.3s, and a radiation thermal heat insulation (RHTI) RHTI 24 of 21.0s and RHTI 24-12 of 7.4s.

The main difference between suit A and B was the clothing ensembles and its protection level. Suit A (traditional turnout gear) contains a trouser and jacket which both need to be worn during firefighting work with protection level 2 according to the standard EN469 (EN 469:2005/ A1:2006 Protective clothing for firefighters – Performance requirements for protective clothing for firefighting). Suit B (new protective clothing concept) contains a coverall with jacket of which the jacket only is mandatory during indoor fire-extinguishing work to provide protection level 2 of the standard EN469. The coverall without the jacket provided only protection level 1. The difference in design of the clothing ensemble (amount of layers) was expected to influence the ventilation and release of heat and so the rise in Tc during the performance of the work.

## 2.3. Study design

The protocol contained the following three stages: (1) concurrent validation measurement; (2) performance of realistic firefighting simulation tasks; (3) concurrent validation measurement. In the concurrent validation measurements in stage 1 and 3, the Tc of the subjects was measured five times at rest with a frequency of one measurement per minute. The Tc was recorded using the Cosinussᵒ (in one ear) and using the CorTemp® and an inner-ear IR thermometer (in other ear) as references. These concurrent validation measurement were performed in a room with a constant Ta of 20.0 2.0 C and RH of 45.0 5.0%. Stage 2 contained a simulation of two realistic firefighting simulations tasks of approximately 15 min per task. First, a selection of standardized tasks were performed selected from the Periodic Preventive Medical Examination (PPMO) protocol namely:

- track including rolling out and up a fire hose of 15 m;

- climbing and descending a ladder (96 steps) with a fire hose of 20 kg over the shoulder;

- crawling through two tunnels (tunnel of 3 m long and 1.2 m height with 3 m between the two tunnels) with a fire hose;

- a demolition operation where a ball of 5 kg needs to be hit the upper side of a basket ten times with a stick of 6 kg on a height of 2.5 m;

- and punching a door with a forcible entry tool of 16 kg.

The PPMO was performed in a room with a constant Ta of 20.0 2.0 C and RH of 45.0 5.0%. Secondly, a hot fire-extinguishing task while wearing self-contained breathing apparatus (SCBA) including:

- extinguishing a fire,

- searching for victim,

- and kneeling in front of a fire of 220 a 225 C.

The Tc of the subjects was recorded continuously using the Cosinussᵒ (one measurement per second) and using the CorTemp® (one measurement per 10 s) as reference. Due to the non-wearable character of the inner-ear thermometer this thermometer could not be included in stage 2. The hot fire-extinguishing task was performed in a practice building with a fireplace of 220 a 225 C. The task started outdoor with a Ta of 13.0 2.0 C, RH of 72.0 5.0% and mean wind speed of 3.2 m/s. The estimated heat radiations was about 4 kW/m2.

لباس محافظ شخصی

دو نوع لباس محافظ شخصی استفاده شد. کت و شلوار A و B. کت و شلوار A یک لباس چرخش سنتی است که از یک شلوار با ژاکت تشکیل شده است (شکل 1 را ببینید). هر دو شلوار و ژاکت حاوی سه لایه بودند. (1) پارچه بیرونی ساخته شده از XT5 Nomex® Delta T؛ (2) سد رطوبت ساخته شده از سلاح های آتش نشانی MO3 Gore-tex®؛ و (3) مانع حرارتی و لایه داخلی ساخته شده از Q01 حرارتی حساس به ویسکوز Nomex® (استاندارد EN469: 2005) (بریستول یکنواخت ، 2007). این کت و شلوار دارای مقاومت بخار آب 28m2Pa / W ، عایق حرارتی حرارتی (HTI) مقاومت HTI 24 از 21.3s و HTI 24-12 از 6.3s و عایق حرارتی تابش گرمائی (RHTI) RHTI 24 از 26.2 ثانیه و RHTI است 24-12 از 7.1s. کت و شلوار B یک مفهوم لباس محافظ جدید است که از یک روپوش با ژاکت تشکیل شده است (به شکل 1 راست مراجعه کنید). ژاکت حاوی دو لایه بود. (1) پارچه بیرونی ساخته شده از TenCate

Millenia ™ Mi 9200 و (2) یک لایه داخلی از TenCate Defender ™ CZ 760 ساخته شده است. این پوشش شامل سه لایه است. (1) پارچه بیرونی ساخته شده از TenCate Mllenia ™ MI 9200 و Tecasafe® Plus XL 9700؛ (2) یک لایه داخلی TenCate Defender ™ CZ760 (قسمت فوقانی) و DM9180 (قسمت پایین)؛ و (3) مانع حرارتی غشای حرارتی TenCate CX 140 (مسترهای ایمنی ، 2017). این کت و شلوار دارای مقاومت بخار آب 12.6m2Pa / W ، عایق حرارتی حرارتی (HTI) مقاومت HTI 24 از 17.4s و HTI 24-12 از 5.3s ، و عایق حرارتی تابش گرمائی (RHTI) RHTI 24 از 21.0s و RHTI 24-12 از 7.4s.

تفاوت اصلی بین کت و شلوار A و B در مجموعه لباس و میزان محافظت از آن بود. کت و شلوار A (چرخ دستی سنتی) شامل یک شلوار و ژاکت است که هر دو در حین کار آتش نشانی با سطح حفاظت 2 مطابق استاندارد EN469 باید پوشیده شوند (EN 469: 2005 / A1: 2006 لباس های حفاظتی برای آتش نشانان - الزامات عملکرد لباس های محافظ برای آتش نشانی). کت و شلوار B (مفهوم لباس محافظ جدید) حاوی یک روپوش با ژاکت است که ژاکت فقط در هنگام خاموش کردن آتش در محیط داخلی برای ارائه سطح حفاظت 2 استاندارد EN469 ضروری است. کاور بدون ژاکت تنها سطح محافظت از آن را در اختیار شما قرار می دهد. 1. تفاوت در طراحی مجموعه لباس (میزان لایه ها) در تأثیر بر تهویه و رهاسازی گرما و از این رو افزایش سطح Tc در حین انجام کار پیش بینی می شد.

2.3 طراحی مطالعه

پروتکل شامل سه مرحله زیر است: (1) اندازه گیری اعتبار سنجی همزمان؛ (2) عملکرد کارهای شبیه سازی آتش نشانی واقع بینانه. (3) اندازه گیری اعتبار سنجی همزمان. در اندازه گیری های اعتبارسنجی همزمان در مرحله 1 و 3 ، غلظت افراد از افراد 5 بار در حالت استراحت با فرکانس یک اندازه گیری در دقیقه اندازه گیری شد. Tc با استفاده از Cosinussᵒ (در یک گوش) و با استفاده از CorTemp® و دماسنج IR داخلی گوش (در گوش دیگر) به عنوان مراجع ثبت شد. این اندازه گیری های اعتبار سنجی همزمان در یک اتاق با TA ثابت از 20.0 2.0 C و RH از 45.0 5.0٪ انجام شد. مرحله 2 شامل یک شبیه سازی از دو کار شبیه سازی آتش نشانی واقع بینانه تقریبا 15 دقیقه در هر کار است. در مرحله اول ، مجموعه ای از کارهای استاندارد شده از پروتکل معاینه پزشکی پیشگیرانه دوره ای (PPMO) انتخاب شده است:

• پیگیری شامل شلنگ و شلنگ آتش 15 متر؛

• بالا رفتن و پایین آمدن از نردبان (96 پله) با شیلنگ آتش 20 کیلوگرم بالای شانه.

• خزیدن از طریق دو تونل (تونل به طول 3 متر و ارتفاع 1.2 متر با 3 متر بین دو تونل) با شیلنگ آتش سوزی.

• عملیات تخریب که در آن توپ 5 کیلوگرم نیاز به ده ضلع بالای سبد ده بار با چوب 6 کیلوگرم در ارتفاع 2.5 متر دارد.

• و مشت زدن درب با ابزار ورود اجباری 16 کیلوگرم.

PPMO در اتاقی با TA ثابت از 20.0 2.0 C و RH 45.0 5.0٪ انجام شد. در مرحله دوم ، یک کار آتش نشانی گرم هنگام پوشیدن دستگاه تنفس خودمختار (SCBA) شامل موارد زیر است:

• خاموش کردن آتش

• جستجوی قربانی ،

• و در مقابل آتش 220 و 225 درجه سانتیگراد زانو بزنید.

Tc از افراد به طور مداوم با استفاده از Cosinussᵒ (یک اندازه گیری در ثانیه) و با استفاده از CorTemp® (یک اندازه گیری در هر 10 ثانیه) به عنوان مرجع ثبت شد. با توجه به ویژگی غیرقابل پوشیدن دماسنج گوش داخلی ، این دماسنج نمی تواند در مرحله 2 باشد. کار آتش نشانی گرم در یک ساختمان تمرینی با شومینه 220 و 225 درجه سانتیگراد انجام شد. T از 13.0 2.0 C ، RH از 72.0 5.0٪ و میانگین سرعت باد 3.2 متر بر ثانیه. تابش حرارتی تخمین زده شده در حدود 4 کیلو وات در متر مربع بود.

In all three stages the subjects wore suit A (trouser and jacket) or suit B (coverall with jacket). After the first round of stage 2 the subject changed suit and performed this stage again in the second suit. To avoid order effects, the order in which the suits were worn alternated per subject. Both measurements in the different suits were performed on the same day in the same order: a concurrent validation measurement, the

PPMO, the hot fire-extinguishing task and a concurrent validation measurement. The PPMO test was done before the hot fire-extinguishing task, because according to regulations after the hot fire-extinguishing task the subjects needed to clean their clothing. Between the three stages, the two tasks and between the measurements in suit A and B, the subjects had time to acclimatize or cool down by passive sitting and drinking water for a period of 10 min. In Table 1, this study design is presented.

To test the in-vivo validity and reliability of the Cosinuss in rest (aim 1), in the concurrent validity study of stage 1 and 3 the Cosinuss was compared to the references CorTemp® and an inner-ear IR thermometer. To test the in-vivo validity and reliability of the Cosinuss during work (aim 2), in stage 2 the Cosinuss was compared to the reference CorTemp®. To explore the change in individual Tc, Tcli and RH of the subjects (aim 3), during stage 2 the Tc was continuously recorded with the Cosinussᵒ and CorTemp® and Tcli and RH were continuously recorded using the ambient conditions box (measurement frequency of one measurement per second). To compare the change in Tc, Tcli and RH in the two types of protective clothing, the two tasks in stage 2 were performed twice; once in suit A and once in suit B.

## 2.4. Data analysis

Field calibration to correct the Cosinuss was explored and conducted based on both the CorTemp® and the inner-ear IR thermometer during stage 1. To correct the Cosinuss the second measurement of stage 1 was used; per subject in one ear the Tc was recorded with the Cosinuss and compared to the Tc recorded with the CorTemp® and the inner-ear IR thermometer in the other ear. The difference between the measured Tc of the Cosinuss and CorTemp® or inner-ear thermometer was considered as the individual correction factor of the Cosinuss. The CorTemp® sensors were factory calibrated and the HQ Inc data loggers were calibrated according to the user instructions.

To test the aims, statistical analysis was performed using IBM SPSS Statistic 25. To test the validity and reliability of the Cosinuss in rest (aim 1), of stage 1 and 3 the fourth concurrent validation measurement was used, including the Tc recorded with the Cosinuss, CorTemp® and inner-ear thermometer. To test the validity and reliability of the Cosinuss during work (aim 2), of stage 2 the mean Tc per task and of both tasks recorded with the Cosinuss and CorTemp® was used. To compare the development of Tc while working in two types of protective clothing (aim 3per subject two datasets were generated, one in suit A and one in suit B. Of stage 2 the mean Tc, Tcli and RH per task and of both tasks per suit was used. Sensitivity analysis was performed to test differences between the fourth and fifth measurement (stage 1 and 3) and the mean of all measurements (all stages), in addition to being performed on only complete datasets (all stages). These sensitivity analyses were performed to verify if the fourth measurement and incomplete datasets are representative

Parametric data were analyzed using the paired t-test and by calculating the intraclass correlation coefficient (ICC, two-way random model). Non-parametric data were analyzed with the Wilcoxon signed rank test. The results are shown with mean or the mean difference (MD) and standard deviation (SD) (mean SD). P-values 0.05 were considered statistically significant. The interpretation of the ICC: ICC<0.39 is poor, 0.40 >ICC>0.59 is moderate, 0.60 >ICC>0.79 is good and ICC0.80 is excellent. The Limits of Agreement (LoA), calculated as 1.96*SDdifference, has an acceptable level of LoA0.50 (). To illustrate if the magnitude of the difference was related to the mean performance, Bland-Altman plots were made (). The individual difference were plotted against the individual mean of the stages ().

در هر سه مرحله ، افراد کت و شلوار A (شلوار و ژاکت) یا کت و شلوار B (کاور با ژاکت) را پوشیدند. بعد از دور اول مرحله 2 موضوع تغییر کرد و این مرحله را دوباره در لباس دوم انجام داد. برای جلوگیری از جلوه های ترتیب ، ترتیب ترتیب لباس در هر موضوع متناوب است. اندازه گیری اعتبارسنجی همزمان ، هر دو اندازه گیری در لباس های مختلف در همان روز در همان روز انجام شدند

PPMO ، وظیفه آتش نشانی گرم و اندازه گیری اعتبار سنجی همزمان. آزمایش PPMO قبل از انجام کار آتش نشانی داغ انجام می شد ، زیرا طبق مقررات پس از انجام وظیفه آتش نشانی داغ ، افراد مورد نیاز برای تمیز کردن لباس خود بودند. بین سه مرحله ، دو وظیفه و بین اندازه گیری های کت و شلوار A و B ، افراد زمان برای نشستن غیرفعال و خنک شدن با نشستن غیرفعال و نوشیدن آب برای مدت زمان 10 دقیقه داشتند. در جدول 1 ، این طرح مطالعه ارائه شده است.

برای آزمایش روایی و پایایی داخل بدن (Cosvinuss) در حالت استراحت (هدف 1) ، در مطالعه روایی همزمان مرحله 1 و 3 ، Cosinuss با منابع CorTemp® و دماسنج IR داخلی گوش مقایسه شد. برای آزمون روایی و روایی درون زراعی Cosinuss در حین کار (هدف 2) ، در مرحله 2 Cosinuss با مرجع CorTemp® مقایسه شد. برای کشف تغییر در Tc ، Tcli و RH در افراد (هدف 3) ، در مرحله 2 Tc به طور مداوم با Cosinussᵒ و CorTemp® ثبت شد و Tcli و RH به طور مداوم با استفاده از جعبه شرایط محیط (فرکانس اندازه گیری یک) ثبت شد. اندازه گیری در ثانیه) برای مقایسه تغییر Tc ، Tcli و RH در دو نوع لباس محافظ ، دو مرحله در مرحله 2 دو بار انجام شد. یک بار در کت و شلوار A و یک بار در کت و شلوار B.

2.4 تحلیل داده ها

کالیبراسیون میدانی برای تصحیح Cosinuss بر اساس CorTemp ther و دماسنج IR داخلی گوش در طی مرحله 1 مورد کاوش و انجام قرار گرفت. برای تصحیح Cosinuss از اندازه گیری دوم مرحله 1 استفاده شد. در هر مورد در یک گوش ، Tc با Cosinuss ثبت شد و با Tc ثبت شده با CorTemp® و دماسنج IR داخلی گوش در گوش دیگر مقایسه شد. تفاوت بین اندازه گیری اندازه گیری شده از Cosinuss و CorTemp® یا دماسنج گوش داخلی به عنوان فاکتور تصحیح فردی Cosinuss در نظر گرفته شده است. سنسورهای CorTemp® کالیبره شدند و گزارشگرهای داده HQ Inc طبق دستورالعمل کاربر کالیبره شدند.

برای آزمون اهداف ، از تجزیه و تحلیل آماری با استفاده از آمار IBM SPSS آمار 25 استفاده شد. برای آزمون روایی و پایایی Cosinuss در حالت استراحت (هدف 1) ، از مرحله 1 و 3 از چهارمین سنجش اعتبار همزمان استفاده شد ، از جمله Tc ثبت شده با Cosinuss ، CorTemp® و دماسنج گوش داخلی. برای آزمون روایی و پایایی Cosinuss در حین کار (هدف 2) ، از مرحله 2 میانگین Tc در هر کار و هر دو کار ضبط شده با Cosinuss و CorTemp® استفاده شد. برای مقایسه توسعه Tc در حین کار در دو نوع لباس محافظ (هدف 3per موضوع از دو مجموعه داده تولید شد ، یکی در لباس A و دیگری در لباس B. از مرحله 2 میانگین Tc ، Tcli و RH در هر کار و هر دو کار در هر مورد استفاده قرار گرفت ، از آنالیز حساسیت برای آزمایش تفاوت بین اندازه گیری چهارم و پنجم (مرحله 1 و 3) و میانگین کلیه اندازه گیری ها (کلیه مراحل) استفاده شد ، علاوه بر این که فقط در مجموعه داده های کامل (کلیه مراحل) انجام می شود. تجزیه و تحلیل برای بررسی اینکه آیا اندازه گیری چهارم و مجموعه داده های ناقص نماینده هستند انجام شد

داده های پارامتری با استفاده از آزمون t زوجی و با محاسبه ضریب همبستگی داخل سلولی (ICC ، مدل تصادفی دو طرفه) تجزیه و تحلیل شدند. دادههای غیرپارامتری با آزمون رتبه امضا شده Wilcoxon مورد تجزیه و تحلیل قرار گرفت. نتایج با میانگین یا میانگین اختلاف (MD) و انحراف استاندارد (SD) (میانگین SD) نشان داده شده است. P-0.05 از نظر آماری معنی دار در نظر گرفته شد. تعبیر ICC: ICC <0.39 ضعیف است ، 0.40> ICC> 0.59 متوسط ​​، 0.60> ICC> 0.79 خوب و ICC0.80 عالی است. محدودیت های توافق نامه (LoA) ، به عنوان 1.96 * SDdifference محاسبه می شود ، سطح قابل قبولی از LoA0.50 () دارد. برای نشان دادن اینکه آیا میزان تفاوت با میانگین عملکرد مربوط است ، توطئه های Bland-Altman () ساخته شد. اختلاف فردی در برابر میانگین فردی مراحل () رسم شد.

3. نتایج

# 3. Results

Eleven firefighters (10 male and one female) with a mean age of 40.1 8.0 years participated in this study. One subject was not able to perform the study in the suit A and due to an error not all data of the Cosinussᵒ and ambient condition box datasets were stored. The usability and sample size of the incomplete datasets varies per aim. Per aim, table and figure the sample size is mentioned. Statistical analysis was performed on as well all available data and only complete datasets.

## 3.1. In-vivo validity and reliability in rest

During the concurrent validation measurements (stage 1 and 3) (n ¼11), the mean Tc recorded with Cosinussᵒ was 36.0 0.8ᵒC with a mean SD within subjects of 0.1 0.1ᵒC. The mean Tc recorded with CorTemp® was 37.5 0.4ᵒC with a mean SD within subjects of 0.1 0.2ᵒC. The mean of Tc recorded with the inner-ear IR was 36.6 0.4 C with a mean SD within subjects of 0.1 0.1 C. Individual correction (in stage 1) based on the CorTemp® resulted in an average correction factor of 1.5 0.7ᵒC and based on the inner-ear IR of 0.6 0.6ᵒC. In Table 2 the MD of the Tc recorded with the Cosinuss and compared to CorTemp® and inner-ear IR are shown.

In stage 1, significant acceptable differences were found between the Cosinuss corrected and compared to CorTemp® (MD ¼0.1 0.1, CI

### Table 2

[0.04; 0.18], p ¼0.006) and the Cosinuss corrected and compared to inner-ear IR (MD ¼0.2 0.2, CI [0.11; 0.32], p ¼0.001). The other combinations showed significantly high differences between the Cosinuss and CorTem®. In stage 3, significantly high differences were found in Tc recorded with the Cosinuss and CorTemp® (MD   0.8, p 0.013). The Cosinuss corrected and compared to CorTemp® showed an acceptable, but non-significant mean difference with a high

SD (MD ¼0.0 1.0, CI [-0.67; 0.65], p ¼0.976). In Table 3 is shown the ICC for Cosinussᵒ, CorTemp® and inner-ear IR.

In stage 1 and 3, without individual correction of the Cosinuss no correlation between the Cosinuss and CorTemp® was observed and a moderate correlation was observed between Cosinuss and inner-ear with an unacceptably high LoA (ICC ¼0.40, p ¼0.044, LoA ¼1.20). The Cosinuss corrected with the CorTemp® (stage 1) resulted in an excellent correlation with an acceptable LoA compared to the CorTemp® (ICC ¼0.94, p ¼0.000, LoA ¼0.20) and a poor correlation compared to the inner-ear IR (ICC ¼0.19, p ¼0.007, LoA ¼0.68). The Cosinuss corrected with the inner-ear IR resulted in a poor correlation compared to the CorTemp® (ICC ¼0.30, p ¼0.005, LoA ¼0.61) and an excellent correlation with an acceptable LoA compared to the inner- ear IR (ICC ¼0.84, p ¼0.000, LoA ¼0.30). No correlation or a poor correlation was found between the CorTemp® and inner-ear IR thermometer (ICC0.24). In Fig. 2, the Bland-Altman plots before working (stage 1) are shown.

After working (stage 3), all correlations between the Cosinuss, CorTemp® and inner-ear IR were lowered to non-significant and poor with unacceptable LoA (ICC0.38, p 0.091, LoA1.71). In Fig. 3, the Bland-Altman plots after working (stage 3) are shown. Sensitivity analysis revealed similar results.

## 3.2. In-vivo validity and reliability during realistic firefighting simulation tasks

During the two realistic firefighting simulation tasks (stage 2) in both suits, the mean Tc recorded with Cosinussᵒ was 36.1 1.1ᵒC (n ¼17). The mean Tc recorded with CorTemp® was 37.6 0.5ᵒC. In Table 4 the MD of the Tc recorded with the Cosinuss and compared to CorTemp® are shown.

In stage 2, the Tc recorded with the non-corrected Cosinuss differs significantly compared to the CorTemp® (MD   0.7, p 0.002). The Cosinuss corrected and compared to the CorTemp® showed an acceptable, but non-significant difference (MD 0.2, p 0.534). In Table 5 is shown the ICC for Cosinussᵒ and CorTemp®.

During the realistic firefighting simulation tasks in stage 2, no correlations between the Cosinuss and the CorTemp® were proven (ICC0.26, p 0.253) with unacceptable high LoA’s (LoA2.20). Sensitivity analysis revealed similar results. In Fig. 4, the Bland-Altman plots during the both realistic firefighting simulation tasks are shown.

During two during realistic firefighting simulation tasks, every

Mean difference (MD) in core temperatur (Tc) (ᵒC) measurements in both suits (n ¼11) of Cosinussᵒ C-med versus reference thermometers; CorTemp® and inner-ear infrared (IR).

Fig. 1. Left: Suit A, traditional turnout gear of Bristol Uniforms containing a trouser with jacket; Right: Suit B, a new protective clothing concept of Safety Masters

containing a coverall combined with a separate outer jacket.

. نتایج

یازده آتش نشان (10 مرد و یک زن) با میانگین سنی 40.1 8.0 سال در این مطالعه شرکت کردند. یک موضوع قادر به انجام مطالعه در کت و شلوار A نبود و به دلیل خطا ، تمام داده های مجموعه داده های Cosinussᵒ و محیط جعبه محیط ذخیره نمی شدند. قابلیت استفاده و اندازه نمونه مجموعه داده های ناقص در هر هدف متفاوت است. در هر هدف ، جدول و شکل اندازه نمونه ذکر شده است. تجزیه و تحلیل آماری و همچنین داده های موجود و فقط مجموعه داده های کامل انجام شد.

3.1 اعتبار داخل و اعتبار در استراحت

در حین اندازه گیری اعتبارسنجی همزمان (مرحله 1 و 3) (11 n n) ، میانگین Tc ثبت شده با Cosinussᵒ 0/0 ᵒ 8/0 درجه سانتیگراد با متوسط ​​SD در افراد 1/0 1/0 بود. میانگین Tc ثبت شده با CorTemp® 4/5 درجه سانتیگراد با میانگین SD در افراد 1/0 ᵒ 2/0 سانتیگراد بود. ميانگين غلظت Tc ثبت شده با گوش داخلي 0.4 36 6/36 درجه سانتيگراد با ميانگين SD در افراد 1/0 C. 1 درجه سانتي گراد بود. در IR داخلی گوش از 0.6 0.6 درجه سانتیگراد. در جدول 2 MD Tc با Cosinuss ثبت شده و در مقایسه با CorTempem و IR داخلی گوش نشان داده شده است.

در مرحله 1 ، تفاوت قابل قبول قابل توجهی بین Cosinuss اصلاح و در مقایسه با CorTemp® (MD ¼ 1/1 0.1 1 ، CI) مشاهده شد.

جدول 2

[0.04؛ 0.18] ، ص .000.006) و Cosinuss با IR داخلی گوش (MD -0.2 0.2 ، CI [0.11 ؛ 0.32] ، P -0.001 اصلاح و مقایسه کردند). ترکیبهای دیگر تفاوتهای زیادی بین Cosinuss و CorTem® نشان دادند. در مرحله 3 ، اختلاف معنی داری زیاد در Tc ثبت شده با Cosinuss و CorTemp® مشاهده شد (MD 0.8 ، P 0.013). Cosinuss تصحیح شده و در مقایسه با CorTemp® تفاوت معنی داری قابل قبول اما غیر معنادار با زیاد نشان داد

SD (MD ¼0.0 1.0 ، CI [-0.67؛ 0.65] ، p -0.976). در جدول 3 ICC برای Cosinussᵒ ، CorTemp® و IR داخلی گوش نشان داده شده است.

در مرحله 1 و 3 ، بدون تصحیح فردی Cosinuss ، هیچ گونه همبستگی بین Cosinuss و CorTemp observed مشاهده نشد و یک همبستگی متوسط ​​بین Cosinuss و گوش داخلی با LoA غیرقابل قبول بالا مشاهده شد (ICC -0.40 ، p -0.044 ، LoA .20 1.20). Cosinuss با CorTemp® (مرحله 1) اصلاح شده منجر به همبستگی بسیار خوبی با یک LoA قابل قبول در مقایسه با CorTemp® (ICC -0.94 ، P -0.000 ، غلظت .0000.000 ، LoA -0.20) و همبستگی ضعیف در مقایسه با داخلی- IR گوش (ICC .10.19 ، ص .000.007 ، LoA -0.68). Cosinuss اصلاح شده با IR داخلی گوش منجر به همبستگی ضعیفی در مقایسه با CorTemp® (ICC -0.30 ، p -0.005 ، LoA -0.61) و ارتباط بسیار خوبی با یک LoA قابل قبول در مقایسه با IR داخلی گوش (ICC -0.84 ، ص ¼0.000 ، LoA. 0.30). هیچ ارتباطی یا همبستگی ضعیفی بین دماسنج IR CorTemp® و داخلی گوش (ICC0.24) مشاهده نشد. در شکل 2 نمودارهای Bland-Altman قبل از کار (مرحله 1) نشان داده شده است.

پس از کار (مرحله 3) ، تمام ارتباطات بین Cosinuss ، CorTemp® و IR داخلی گوش به غیر مهم و فقیر با LoA غیرقابل قبول کاهش یافت (ICC0.38 ، P 0.091 ، LoA1.71). در شکل 3 نمودارهای Bland-Altman پس از کار (مرحله 3) نشان داده شده است. تجزیه و تحلیل حساسیت نتایج مشابهی را نشان داد.

3.2 اعتبار و اعتبار در داخل بدن در حین انجام کارهای شبیه سازی آتش نشانی واقعی

در طول دو کار شبیه سازی آتش نشانی واقع بینانه (مرحله 2) در هر دو لباس ، میانگین Tc ثبت شده با Cosinussᵒ 1/36 درجه سانتیگراد (17 n N) بود. ميانگين غلظت Tc ثبت شده با CorTemp® 5/0 6 / 0CC بود. در جدول 4 MD از Tc ثبت شده با Cosinuss و در مقایسه با CorTemp® نشان داده شده است.

در مرحله 2 ، Tc ثبت شده با Cosinuss اصلاح نشده در مقایسه با CorTemp® تفاوت معنی داری دارد (MD 0.7 ، 0.002 p). Cosinuss با CorTemp® اصلاح و در مقایسه با CorTemp® تفاوت قابل قبول اما غیر معناداری را نشان داد (MD 0.2 ، p 0.534). در جدول 5 ICC برای Cosinussᵒ و CorTemp® نشان داده شده است.

در طی کارهای شبیه سازی آتش نشانی واقع بینانه در مرحله 2 ، هیچ ارتباطی بین Cosinuss و CorTemp® اثبات نشده است (ICC0.26 ، 0.253 ص) با LoA زیاد غیر قابل قبول (LoA2.20). تجزیه و تحلیل حساسیت نتایج مشابهی را نشان داد. در شکل 4 ، توطئه های Bland-Altman در هر دو کار شبیه سازی آتش نشانی واقع بینانه نشان داده شده است.

در طول دو کار شبیه سازی آتش نشانی واقع بینانه ، هر

میانگین اختلاف (MD) در اندازه گیری درجه حرارت مرکزی (TC) در دمای هر دو لباس (11 n N) Cosinuss Cos C-med در مقابل دماسنجهای مرجع. CorTemp® و مادون قرمز گوش داخلی (IR).

شکل 1. سمت چپ: کت و شلوار A ، لباس مشارکت سنتی لباسهای بریستول حاوی یک شلوار با ژاکت؛ سمت راست: کت و شلوار B ، یک مفهوم لباس محافظ جدید از استادان ایمنی

حاوی یک پوشش کامل با یک ژاکت بیرونی جداگانه.

subject showed a different pattern in the change in Tc, Tcli and RH. The mean change in Tc in both suits recorded with the Cosinuss was 0.06 0.14 C/minute and the mean change in Tc recorded the CorTemp® was 0.02 0.10 C/minutes. The mean Tcli in both suits increased with 0.4 0.4 C/minute with a mean Tcli of 29.7 2.1 C and the RH increased with 1.6 3.2%/minute with a mean RH of 68.4 10.8%. In Table 6 the mean Tc recorded with the Cosinuss and CorTemp® and Tcli and RH per task per suit are shown.

The maximum Tc recorded with the Cosinuss was 40.1 C and with the CorTemp® 38.4 C (both during the fire-extinguishing task). However, on average the maximum Tc recorded with the Cosinuss is significant lower compared to the CorTemp® during both tasks (during PPMO MD ¼1.1 1.2, CI [0.45; 1.68], p ¼0.002, during the fire- extinguishing task MD ¼1.1 1.1, CI [0.51; 1.73], p ¼0.001). In Table 7 the MD and ICC of the Tc recorded with the Cosinuss CorTemp®, Tcli and RH in both suits between the Periodic Preventive Medical Examination (PPMO) and hot fire-extinguishing task are shown.

Fig. 2. Bland-Altman plots of the mean core temperature versus the mean temperature difference before working (stage 1) in both suits between the non-corrected and corrected Cosinuss compared to the CorTemp® and inner-ear IR with mean (black line) and upper and lower Limit of Agreement (LoA) (black dotted line) and zero-line (blue line), n ¼11. (For interpretation of the references to colour in this figure legend, the reader is referred to the Web version of this article.)

correlation was found be between the Tcli of suit A and B (ICC ¼0.74, CI [0.06; 0.96], p ¼0.017, n ¼8). No other significant differences were found between suit A and B with a good to excellent correlation between the Tc recorded with Cosinuss (ICC0.74, p 0.017, n ¼7). Additionally, no significant difference in the Tc, Tcli and RH change between the suits during the tasks was found (p 0.308). Sensitivity analysis revealed similar results.

# 4. Discussion

The Tc of the non-corrected Cosinuss showed no correlations with the CorTemp® and a poor to moderate correlation with inner-ear IR thermometry, indicating the impact of individual correction. Individual correction resulted in excellent correlations. However it should be noted that this depends on the calibration method and it is unknown if

individual calibration using the CorTemp® performed once is valid over longer periods of time. Moreover, no correlation was found between the CorTemp® and inner-ear IR and no correlation or a poor correlation was found between the CorTemp® and Cosinussᵒ during the two tasks, indicating a difference due to the measurement method and location (Towey et al., 2017) (Taylor et al., 2014). The measured change in Tc by the Cosinussᵒ, as well as by inner-ear IR, is expected to be the local temperature and/or caused by a preliminary or side effect (Levander and Grodzinsky, 2017) (Kuht and Farmery, 2014). Therefor calibration should be done using at least an invasive reference, such as CorTemp®, however calibration with a controlled thermostat bath at different temperatures is preferred. The correlations were poor after performance of the tasks. This is probably caused by a secondary human-device interaction factor, e.g., a shift of the sensor in the ear caused by movement. Overall, the Cosinuss is an invalid method for measuring core temperature of firefighters during the performance of their job. The influence of the consumption of hot and cold food and drinks was noted by the research and clearly visible in the results of the CorTemp® (Collin et al., 2015).

The strength of this study is that validation of the Tc measurements was performed using the invasive      as well as the inner-ear clinical standard (Moran-Nabarro et al., 2018) (Nederlands Huisartsen Genootschap, 2016) (Itani et al., 2018) (Ouahrani et al., 2017) (Nadipi Reddy et al., 2017). Furthermore this applied research was performed in a field situation with realistic firefighting tasks which are used to train and test firefighters (Havenith and Heus, 2004). Also, in addition to insight in the validity of the wearable thermometer, research was performed on two different types of protective suits and the difference in the development of heat stress between tasks was assessed. The main limitation was the duration of the tasks. The schedule was planned 30 min per task, but participants only took approximately 10 up to 15 min per task. This time span will result in only limited heating up of the body and less useful results. Especially in the case of exploring differences in heat development in the two types of protective clothing 10 min was not long enough to gain realistic insights and draw conclusions which are representing real-life firefighting tasks in protective

موضوع الگوی متفاوتی در تغییر در Tc ، Tcli و RH نشان داد. میانگین تغییر در Tc در هر دو لباس ثبت شده با Cosinuss 0/0 0. 14/0 در دقیقه و میانگین تغییر در Tc ثبت شده در CorTemp® 0/0 2 10/0 در دقایقی بود. میانگین Tcli در هر دو لباس با 0.4 0.4 درجه سانتیگراد در دقیقه با میانگین Tcli 29.7 2.1 C افزایش یافته و RH با 1.6 3.2٪ در دقیقه با میانگین RH از 68.4 10.8٪ افزایش یافته است. در جدول 6 میانگین Tc ثبت شده با Cosinuss و CorTemp® و Tcli و RH در هر کار برای هر کت و شلوار نشان داده شده است.

حداکثر Tc ثبت شده با Cosinuss برابر 40.1 درجه سانتیگراد و با CorTemp® 38.4 C (هر دو در حین کار اطفای حریق) بود. با این حال ، به طور متوسط ​​حداکثر Tc ثبت شده با Cosinuss در مقایسه با CorTemp® در هر دو کار (در طول PPMO MD-1.1 1.2 ، CI [0.45 ؛ 1.68] ، P -0.002 ، در طول کار اطفای حریق قابل توجهی پایین تر است). ¼1.1 1.1، CI [0.51؛ 1.73]، p ¼0.001). در جدول 7 MD و ICC از Tc ثبت شده با Cosinuss CorTemp® ، Tcli و RH در هر دو لباس بین معاینه پزشکی پیشگیری دوره ای (PPMO) و کار خاموش کردن آتش نشان داده شده است.

شکل 2. توطئه Bland-Altman از میانگین دمای هسته در مقابل اختلاف میانگین درجه حرارت قبل از کار (مرحله 1) در هر دو لباس بین Cosinuss اصلاح نشده و اصلاح شده در مقایسه با CorTemp® و IR داخلی گوش با میانگین (خط سیاه) ) و حد بالا و پایین توافق نامه (LoA) (خط خال خال) و خط صفر (خط آبی) ، N-11. (برای تفسیر اشارات به رنگ در این افسانه شکل ، خواننده به نسخه وب این مقاله مراجعه می کند.)

همبستگی بین Tcli کت و شلوار A و B وجود دارد (ICC -0.74 ، CI [0.06 ؛ 0.96] ، P-01.017 ، N-8). تفاوت معنی داری دیگر بین کت و شلوار A و B با همبستگی خوب و عالی بین Tc ثبت شده با Cosinuss مشاهده نشد (ICC0.74 ، 0.017 ، P-7). علاوه بر این ، تفاوت معنی داری در تغییر Tc ، Tcli و RH بین لباس در حین انجام کار وجود نداشت (308/0 = P). تجزیه و تحلیل حساسیت نتایج مشابهی را نشان داد.

4. بحث

TC از Cosinuss اصلاح نشده هیچ ارتباطی با CorTemp® و همبستگی ضعیف تا متوسط ​​با دماسنج IR IR داخلی گوش نشان نداد ، که نشان دهنده تأثیر تصحیح فرد است. اصلاح فردی منجر به همبستگی بسیار خوبی شد. اما لازم به ذکر است که این به روش کالیبراسیون بستگی دارد و مشخص نیست که آیا

کالیبراسیون فردی با استفاده از CorTemp® انجام شده یک بار در مدت زمان طولانی تری معتبر است. علاوه بر این ، هیچ ارتباطی بین CorTemp® و IR داخلی گوش مشاهده نشد و هیچ ارتباطی یا همبستگی ضعیفی بین CorTemp® و Cosinussᵒ در طول دو کار یافت نشد ، که نشان دهنده اختلاف به دلیل روش اندازه گیری و موقعیت مکانی است (تووی و همکاران). ، 2017) (تیلور و همکاران ، 2014). انتظار می رود تغییر اندازه گیری Tc توسط Cosinussᵒ ، و همچنین توسط IR داخلی گوش ، دمای محلی و / یا ناشی از اثر اولیه یا جانبی باشد (Levander and Grodzinsky، 2017) (Kuht and Farmery، 2014). از این رو کالیبراسیون باید با استفاده از حداقل یک مرجع تهاجمی ، مانند CorTemp® انجام شود ، اما کالیبراسیون با حمام ترموستات کنترل شده در دماهای مختلف ترجیح داده می شود. همبستگی ها پس از انجام کارها ضعیف بودند. این احتمالاً به دلیل وجود یک تعامل ثانویه دستگاه بین انسان ، به عنوان مثال ، تغییر حسگر در گوش ناشی از حرکت ایجاد شده است. به طور کلی ، Cosinuss یک روش نامعتبر برای اندازه گیری دمای هسته آتش نشانان در حین انجام کار خود است. تأثیر مصرف مواد غذایی و نوشیدنی های گرم و سرد توسط این تحقیقات مشخص شد و در نتایج CorTemp® به وضوح قابل مشاهده است (کالین و همکاران ، 2015).

قوت این مطالعه این است که اعتبارسنجی اندازه گیری های Tc با استفاده از استاندارد کلینیکی تهاجمی و همچنین گوش داخلی انجام می شود (موران-نبارو و همکاران ، 2018) (Nederlands Huisartsen Genootschap، 2016) (Itani et al.، 2018 ) (Ouahrani et al.، 2017) (Nadipi Reddy et al.، 2017). علاوه بر این ، این تحقیق کاربردی در یک موقعیت میدانی با کارهای آتش نشانی واقع بینانه انجام می شود که برای آموزش و آزمایش آتش نشانان مورد استفاده قرار می گیرند (Havenith and Heus، 2004). همچنین ، علاوه بر بینش در اعتبار دماسنج پوشیدنی ، تحقیق در مورد دو نوع لباس های محافظ مختلف انجام شد و تفاوت در ایجاد تنش گرما بین وظایف مورد بررسی قرار گرفت. محدودیت اصلی مدت کار بود. برنامه زماني 30 دقيقه براي هر كار طراحي شده بود ، اما شركت كنندگان فقط تقريبا 10 تا 15 دقيقه در هر كار انجام دادند. این بازه زمانی منجر به گرم شدن محدود بدن و نتایج کمتری خواهد شد. به خصوص در مورد کشف اختلافات در تولید گرما در دو نوع لباس محافظ ، 10 دقیقه به اندازه کافی برای به دست آوردن بینش واقعی و نتیجه گیری که نمایانگر کارهای واقعی آتش نشانی در محافظت هستند ، نیست.

Fig. 4. Bland-Altman plots of the mean core temperature versus the mean temperature difference during both firefighting simulation tasks (stage 2) in both suits between the non-corrected and corrected Cosinuss compared to the CorTemp® with mean (black line) and upper and lower Limit of Agreement (LoA) (black dotted line) and zero-line (blue line), n ¼17. (For interpretation of the references to colour in this figure legend, the reader is referred to the Web version of this article.) Table 6

Mean and maximum core temperature (Tc) (ᵒC) recorded with the CorTemp® and (non-corrected) Cosinussᵒ, temperature inside de clothing (Tcli) (C) and humidity (RH) (%) per (Periodic Preventive Medical Examination (PPMO) and hot fire-extinguishing) task per suit A and B.

PPMO

Fire-extinguishing task

## Table 7

Mean difference in core temperature (Tc) (C) recorded with the (non-corrected) Cosinussᵒ and the CorTemp®, temperature inside the clothing (Tcli) (ᵒC), humidity (RH) and their change (Δ) between the Periodic Preventive Medical Examination (PPMO) and hot fire-extinguishing task in both suits.

## Table 8

Mean difference in core temperature (Tc) (C) recorded with the (non-corrected) Cosinussᵒ and the CorTemp®, temperature inside the clothing (Tcli) (ᵒC), humidity (RH) and their change (Δ) per (Periodic Preventive Medical Examination (PPMO) and hot fire-extinguishing) task between suit A and B.

clothing. Besides, it would have been interested to also include suit B without the jacket (only coverall with protection level 1) to explore the change in Tc, Tcli and RH during the PPMO task and to falsify is this new protective clothing concept is lowering the heat stress development during no firefighting activities. In addition, some datasets were incomplete. This resulted in a lower amount of data which could result in the risk over overestimating the results. In this case it was not a selective drop-out which significantly could have influence the data. Moreover, according to a power analysis and the two measurements (suits A and B) per subject, enough data was gathered for analysis resulting in no need to replace these missing data. Besides, the failures of the systems provided useful input about the usability and weaknesses of the systems. Another limitation could be the difference in measurement frequency between the thermometers, although not much influence is expected since the temperature rise is only gradually.

To be able to apply the Cosinuss, the system should be valid and reliable. The urge for individual calibration should be investigated, as well as using the Cosinuss in a multivariable system to increase the accuracy of the Tc prediction. In addition, to improve the low correlation during and after performance of the tasks (stage 2 and 3), movements of the Cosinussᵒ should be limited. Further research should study the change in Tc over longer periods of time and with more participants. Furthermore, improvements are necessary to fix the device more stable in the ear. Without movement of the sensor relative to the ear its performance might be much improved. And the integration of the ear part with the ambient condition box would possibly improve the usability of this system. By eliminating the chest box, which was not experienced as very comfortable, a multivariable single-instrument system could be created. Above all, to make this system useable and safe in use for firefighters during their work, the system needs to be made heat- and fire-resistant. Currently the Cosinuss is made of a thermoplastic with silicone developed to be used in ambient temperatures of 15 to 55 C (Cosinuss, 2016), which can cause burns during long-term use in heat exposure nor is fulfilling the safety standards.

If the reliability of the Cosinuss can be improved, as part of a multivariable system (Richmond et al., 2015) the Cosinuss could play a role in predicting Tc and heat storage. However, a reliable single-instrument system is preferred. A wearable thermometer could overcome the disadvantages of the temperature pill if it is an accurate, valid and reliable measurement system that does not underestimate the Tc. A wearable thermometer is non-invasive, useable for all body proportions, reusable with low costs and not influenced by food and liquid intake   (Saurabh et al., 2014) (Taylor et al., 2014) (Lim et al., 2008). In addition, it is immediately useable without a waiting time (Collin et al., 2015). Wearable non-invasive thermometers are of interest in multiple fields for (health) monitoring purposes    (Li et al., 2019) (Diaz-Piedra et al., 2019) (Steck et al., 2011). Next to firefighters, other physical active workers are at risk of work-related overheating and are interested in such a device (Pancardo et al., 2015), including astronauts who want to launch wearables into space to monitor their health during their stay (Jones, 2006). However, the Cosinuss in its present form is not suitable for this.

# . توطئه Bland-Altman از میانگین دمای هسته در مقابل اختلاف میانگین دما در هر دو کار شبیه سازی آتش نشانی (مرحله 2) در هر دو لباس بین Cosinuss اصلاح نشده و اصلاح شده در مقایسه با CorTemp® با میانگین (خط سیاه) و بالا و پایین حد توافق نامه (LoA) (خط خال خال) و خط صفر (خط آبی) ، N-17. (برای تفسیر ارجاعات به رنگ در این افسانه شکل ، خواننده به نسخه وب این مقاله ارجاع داده شده است.) جدول 6

# میانگین و حداکثر دمای هسته (TC) (CC) ثبت شده با CorTemp® و (اصلاح نشده) Cosinussᵒ ، دمای داخل لباس (Tcli) (C) و رطوبت (RH) (٪) در هر (آزمایش پزشکی پیشگیرانه دوره ای (PPMO ) و آتش خاموش) کار در هر کت و شلوار A و B

# جدول 7

# میانگین اختلاف دمای هسته (Tc) (C) ثبت شده با (اصلاح نشده) Cosinussᵒ و CorTemp® ، درجه حرارت داخل لباس (Tcli) (ᵒC) ، رطوبت (RH) و تغییر آنها (Δ) بین دوره پیشگیری معاینه پزشکی (PPMO) و کار آتش نشانی گرم در هر دو لباس.

# جدول 8

# میانگین اختلاف در دمای هسته (Tc) (C) ثبت شده با (اصلاح نشده) Cosinussᵒ و CorTemp® ، درجه حرارت داخل لباس (Tcli) (CC) ، رطوبت (RH) و تغییر آنها (Δ) در هر (دوره ای پیشگیرانه) معاینه پزشکی (PPMO) و آتش خاموش داغ) بین کت و شلوار A و B

# تن پوش. علاوه بر این ، علاقه مند خواهد بود که کت و شلوار B بدون ژاکت (فقط کاور با سطح محافظت از سطح 1) را نیز برای کشف تغییر در Tc ، Tcli و RH در حین انجام وظیفه PPMO و جعل این موضوع بدانید که آیا این مفهوم لباس محافظ جدید در حال کاهش حرارت است. توسعه استرس در طول هیچ فعالیت آتش نشانی. علاوه بر این ، برخی از مجموعه داده ها ناقص بودند. این امر منجر به مقدار کمتری از داده ها می شود که می تواند منجر به خطر بیش از حد ارزیابی نتایج شود. در این حالت ، این یک انتخاب ترکیبی نبود که به طور قابل توجهی می تواند بر داده ها تأثیر بگذارد. علاوه بر این ، با توجه به تجزیه و تحلیل قدرت و دو اندازه گیری (مناسب A و B) در هر موضوع ، داده های کافی برای تجزیه و تحلیل جمع آوری شده و در نتیجه نیازی به جایگزینی این داده های مفقوده نیست. علاوه بر این ، عدم موفقیت سیستم ها در مورد قابلیت استفاده و ضعف سیستم ها مفید بوده است. محدودیت دیگر می تواند تفاوت در فرکانس اندازه گیری بین دماسنج ها باشد ، گرچه از زمان افزایش دما فقط به تدریج انتظار نمی رود تأثیر زیادی داشته باشد.

# برای اینکه بتوانید از Cosinuss استفاده کنید ، این سیستم باید معتبر و قابل اعتماد باشد. نیاز به کالیبراسیون فردی باید بررسی شود ، همچنین با استفاده از Cosinuss در یک سیستم چند متغیره برای افزایش دقت پیش بینی Tc. علاوه بر این ، برای بهبود همبستگی کم در طول و بعد از انجام کارها (مرحله 2 و 3) ، حرکات Cosinussᵒ باید محدود باشد. تحقیقات بیشتر باید تغییر در Tc را در دوره های طولانی تر و با شرکت کنندگان بیشتر انجام دهد. علاوه بر این ، اصلاحاتی لازم است که دستگاه در گوش پایدارتر باشد. بدون حرکت سنسور نسبت به گوش عملکرد آن ممکن است بسیار بهبود یابد. و ادغام قسمت گوش با جعبه شرایط محیط احتمالاً قابلیت استفاده از این سیستم را بهبود می بخشد. با از بین بردن جعبه قفسه سینه ، که به عنوان خیلی راحت تجربه نشده بود ، می توان یک سیستم تک ابزار چند متغیره ایجاد کرد. مهمتر از همه ، برای استفاده از این سیستم در استفاده و ایمن بودن آتش نشانان در حین کار ، سیستم باید از نظر گرما و مقاوم در برابر آتش ساخته شود. در حال حاضر Cosinuss از ترموپلاستیک با سیلیکون ساخته شده است که در دمای محیط 15 تا 55 درجه سانتیگراد مورد استفاده قرار می گیرد (Cosinuss، 2016) ، که می تواند در اثر استفاده طولانی مدت در معرض گرما باعث سوختگی شود و همچنین استانداردهای ایمنی را برآورده نمی کند.

# در صورت بهبود قابلیت اطمینان Cosinuss ، به عنوان بخشی از یک سیستم چند متغیره (ریچموند و همکاران ، 2015) ، Cosinuss می تواند در پیش بینی Tc و ذخیره سازی حرارت نقش داشته باشد. با این حال ، یک سیستم تک ابزار قابل اعتماد ترجیح داده می شود. دماسنج پوشیدنی اگر یک سیستم اندازه گیری دقیق ، معتبر و قابل اطمینان باشد که Tc را دست کم نگیرد می تواند بر معایب قرص دما غلبه کند. دماسنج پوشیدنی غیر تهاجمی است ، قابل استفاده برای تمام نسبت های بدن ، قابل استفاده مجدد با هزینه های کم و تحت تأثیر مواد غذایی و مصرف مایع نیست (Saurabh et al.، 2014) (Taylor et al.، 2014) (Lim et al.، 2008) . علاوه بر این ، بلافاصله بدون زمان انتظار قابل استفاده است (کالین و همکاران ، 2015). دماسنجهای پوشیدنی غیر تهاجمی در زمینه های مختلف برای اهداف نظارت بر سلامت (لی و همکاران ، 2019) مورد علاقه هستند (Diaz-Piedra et al.، 2019) (Steck et al.، 2011). در کنار آتش نشانان ، سایر کارگران فعال بدنی نیز در معرض گرمای بیش از حد کار قرار دارند و به چنین وسیله ای علاقه مندند (Pancardo et al.، 2015) از جمله فضانوردانی که می خواهند پوشیدنی را به فضا پرتاب کنند تا در طول مدت اقامت خود بر سلامت خود نظارت کنند (جونز ، 2006). با این حال ، Cosinuss به شکل فعلی خود برای این مناسب نیست.

# 5. Conclusion

The validity of the Cosinuss C-med was not confirmed in this study. Without individual correction, the Cosinuss showed poor to moderate correlations resulting in an invalid, but reliable system. With individual correction, depending on the instruments used for field calibration in resting conditions an excellent correlation for measuring core temperature was found, resulting in a valid and reliable system. During and after subject performance of the tasks, non-significant poor correlations were found. This is most likely caused by firstly the measurement location and secondly non-adequate fixation of the Cosinuss causing a movement of the sensor relative to the ear. This indicates that Cosinuss is an invalid method for measuring core temperature of firefighters during the performance of their job. The temperature inside the suit was significantly higher in the new protective clothing concept compared to the traditional turnout gear. No significant differences were found in the change in core temperature, temperature and humidity inside the suit.

# Funding

This work, part of SPRINT@Work, was supported by the European Regional Development Fund, the province and municipality of Groningen, and the province of Drenthe [grant number T-3036, 2013]. Cosinussᵒ has supplied 6 Cosinussᵒ C-med during this study.

# Declaration of competing interest

The authors declare that they have no known competing financial interests or personal relationships that could have appeared to influence the work reported in this paper.

# Acknowledgements

We thank ‘Veiligheidsregio Groningen’, in specific Bert Hindriks and Arie Smedinga, for the opportunity to perform this study and their infinite support. We appreciate the contribute of Negotica (Groningen, The Netherlands) during the development of the CORTES2/3, and ProCare (Groningen, The Netherlands) for their effort with the temperature pills.

5- نتیجه گیری

روایی اعتبار Cosinuss C-med در این مطالعه تأیید نشده است. بدون تصحیح فردی ، Cosinuss همبستگی ضعیف و متوسط ​​را نشان داد که منجر به یک سیستم نامعتبر ، اما قابل اعتماد می شود. با تصحیح فردی ، بسته به سازهای مورد استفاده برای کالیبراسیون میدانی در شرایط استراحت ، یک همبستگی عالی برای اندازه گیری دمای هسته یافت شد ، که منجر به یک سیستم معتبر و قابل اعتماد می شود. در طی و بعد از اجرای موضوع وظایف ، همبستگی ضعیف و معنادار مشاهده نشد. این امر به احتمال زیاد در مرحله اول محل اندازه گیری و ثانیاً عدم تثبیت مناسب Cosinuss است که باعث ایجاد حسگر نسبت به گوش می شود. این نشان می دهد که Cosinuss یک روش نامعتبر برای اندازه گیری دمای هسته آتش نشانان در حین انجام کار خود است. درجه حرارت داخل کت و شلوار در مفهوم لباس های محافظ جدید نسبت به چرخ دنده های مشارکت سنتی به طور قابل توجهی بالاتر بود. تفاوت معنی داری در تغییر دمای هسته ، دما و رطوبت داخل لباس مشاهده نشد.

منابع مالی

این کار ، بخشی از SPRINT @ Work ، توسط صندوق توسعه منطقه ای اروپا ، استان و شهرداری گرونینگن ، و استان درنته پشتیبانی می شود. [شماره کمک هزینه T-3036 ، 2013]. Cosinussᵒ در طول این مطالعه 6 Cosinussᵒ C-med را تهیه کرده است.

اعلان منافع رقابتی

نویسندگان اظهار داشتند که آنها هیچ علایق مالی و یا روابط شخصی رقابتی شناخته شده ای را که ممکن است به نظر می رسد تأثیرگذار بر کار گزارش شده در این مقاله است ، ندارند.

سپاسگزاریها

ما از "Veiligheidsregio Groningen" به ویژه Bert Hindriks و Arie Smedinga بخاطر فرصتی برای انجام این مطالعه و حمایت بی نهایت آنها تشکر می کنیم. ما از مشارکت Negotica (گرونینگن ، هلند) در توسعه CORTES2 / 3 و ProCare (گرونینگن ، هلند) برای تلاش خود با قرص های دما قدردانی می کنیم.